## Instructions
LipidQC
A community data tool to support lipidomics data workflows and compare experimental concentrations against the consensus means (nmol/mL) from the NIST Lipidomics Interlaboratory Study for lipid species present in NIST SRM 1950 - "Metabolites in Frozen Human Plasma".
Introduction
- | Data should be formatted as a standard data table with lipid species as rows and samples/replicates as columns.
- | Please enable macros and content if necessary.
- | Report measurements for all analytes in (nmol/mL).
- | Submit only one lipid species per entry (no isobaric entries).
See additional details in Supplemental Information below.
- | This tool produces a result table and "landing strip" style control plot comparing experimental results with the NIST interlaboratory consensus means for the chosen material.
- | Any values changed after initial entry will update throughout LipidQC.
Instructions
- | Step 1 | Save a copy of this workbook with a meaningful name (automatic). Overwriting the original file is possible but not recommended.
- | Step 2 | Select the material against which to compare your values.
- | Step 3
- | Step 4
v1.0.20170809
Supplemental Information
Inputs: Compound Name, By total carbons and double bonds
Start with the acronym found in the worksheet for that lipid type
Then add a space, the total carbons and double bonds separated by a colon ":"
For example: MG 18:0, PC 34:1, PE 34:2, TG 62:4, etc.
Inputs: Compound Name, By acyl constituents
Start with the acronym found in the current worksheet (eg. TG, PC, PMe…)
Follow with "(", and acyl chains, each denoted by total carbons and unsaturations
Each acyl chain can be seperated by an "/" (known fatty acyl chain position)
Each acyl chain can be seperated by an "_" (unknown fatty acyl chain position)
For example: PC(18:0/16:0), PE(20:0_22:6), TG(16:0_18:1_20:1)
Inputs: Ether-Linked Lipids (By total carbons and double bonds)
Start with the the acronym found in the worksheet for that lipid type
Then add a space and a "P-" for plasmenyl ethers and "O-" for plasmanyl ethers
Follow with the total carbons and double bonds separated by a colon ":"
For example: MG P-18:0, PC O-34:1, PE P-34:2, TG P-62:4, etc.
Inputs: Ether-Linked Lipids (By acyl constituents)
Start with the acronym found in the current worksheet (eg. TG, PC, PMe…)
Follow with "(", and acyl chains, each denoted by total carbons and unsaturations
The first acyl chain should start with a "P-" for plasmenyl ether-linked lipid species and "O-" for plasmanyl ether-linked lipid species
Each acyl chain can be seperated by an "/" or "_" for known or unknown fatty acyl chain positions
Note that di- or tri-vinyl ether and ether linked lipids are not supported
For example: PC(P-16:0_18:0), PE(O-16:1/18:2), TG(O-16:0_18:1_18:3)
Inputs: Isobaric Species (Not Supported)
Submit only one lipid species per entry.
Accepted: PC O-40:7, PC P-40:6, PC 39:7 (separate entries)
Not Supported: PC O-40:7/PC P-40:6 (single entry)
For public beta release under the NIST Software Disclaimer.
Associated with [publication DOI: XXXXXXXXXXXX]
Please refer questions, bug reports, or suggestions to authors at:
Candice Ulmer, PhD |  |  | candice.ulmer@nist.gov
Jeremy Koelmel |  |  | jeremykoelmel@gmail.com
Jared M. Ragland |  |  | jared.ragland@nist.gov
Christina M. Jones |  |  | christina.jones@noaa.gov | Version 1.0
John Bowden, PhD |  |  | john.bowden@nist.gov | Last Updated 09 August 2017
Enter values
Compare (automatic)

## Enter Values (hidden)
Input Lipid Species | Input Concentration (nmol/mL) Be sure to include appropriate headers for replicates in the row below (e.g., Sample 1, Sample 2, Sample 3)
Lipid Name

## certlist (hidden)
SRM | Title | Certificate Date | Link to pdf | Comment | https://www-s.nist.gov/srmors/certificates/view_certPDF.cfm?certificate= | urines(many) | levels 2 and 3 of 955c
SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | https://www-s.nist.gov/m-srmors/certificates/1950.pdf

## consensus values (hidden)
ID | Material Number | Material Name | DateIssued | Measureand | Key | CertRef | Value | ExpandedUncertainty | 95low | 95high | u | UnitNumerator | UnitDenominator | NumLabs | StandardUncertainty | CoefficientOfDispersion(%) | k
1 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | FFA 16:0 |  | Ref | 43 | 26 | 17 | 69 | 13 | nmol | mL | 5 | 13 | 31 | 2
2 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | FFA 18:3 |  | Ref | 2.9 | 1.24 | 1.66 | 4.14 | 0.62 | nmol | mL | 6 | 0.62 | 21 | 2
3 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | FFA 20:4 |  | Ref | 4.7 | 3 | 1.7 | 7.7 | 1.5 | nmol | mL | 7 | 1.5 | 31 | 2
4 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | FFA 20:5 |  | Ref | 0.42 | 0.112 | 0.308 | 0.532 | 0.056 | nmol | mL | 7 | 0.056 | 13 | 2
5 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | FFA 22:6 |  | Ref | 1.5 | 0.34 | 1.16 | 1.84 | 0.17 | nmol | mL | 8 | 0.17 | 11 | 2
6 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 12-HETE |  | Ref | 0.0068 | 0.003 | 0.0038 | 0.0098 | 0.0015 | nmol | mL | 5 | 0.0015 | 23 | 2
7 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 15-HETE |  | Ref | 0.0024 | 0.00128 | 0.00112 | 0.00368 | 0.00064 | nmol | mL | 5 | 0.00064 | 27 | 2
8 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 5-HETE |  | Ref | 0.01 | 0.0026 | 0.0074 | 0.0126 | 0.0013 | nmol | mL | 5 | 0.0013 | 13 | 2
9 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | DAG 30:0 |  | Ref | 0.83 | 0.34 | 0.49 | 1.17 | 0.17 | nmol | mL | 7 | 0.17 | 20 | 2
10 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | DAG 34:1 |  | Ref | 6.1 | 4.8 | 1.3 | 10.9 | 2.4 | nmol | mL | 16 | 2.4 | 40 | 2
11 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | DAG 36:2 |  | Ref | 6.2 | 4.4 | 1.8 | 10.6 | 2.2 | nmol | mL | 16 | 2.2 | 36 | 2
12 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | DAG 36:3 |  | Ref | 8.4 | 6.6 | 1.8 | 15 | 3.3 | nmol | mL | 15 | 3.3 | 39 | 2
13 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | DAG 36:4 |  | Ref | 2.8 | 2 | 0.8 | 4.8 | 1 | nmol | mL | 12 | 1 | 38 | 2
14 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 46:2 |  | Ref | 3.6 | 2.6 | 1 | 6.2 | 1.3 | nmol | mL | 8 | 1.3 | 37 | 2
15 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 48:0 |  | Ref | 4.5 | 2.4 | 2.1 | 6.9 | 1.2 | nmol | mL | 10 | 1.2 | 26 | 2
16 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 48:1 |  | Ref | 13 | 6.4 | 6.6 | 19.4 | 3.2 | nmol | mL | 16 | 3.2 | 24 | 2
17 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 48:2 |  | Ref | 16 | 5.6 | 10.4 | 21.6 | 2.8 | nmol | mL | 15 | 2.8 | 18 | 2
18 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 48:4 |  | Ref | 1.3 | 0.46 | 0.84 | 1.76 | 0.23 | nmol | mL | 5 | 0.23 | 18 | 2
19 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 49:1 |  | Ref | 2 | 0.84 | 1.16 | 2.84 | 0.42 | nmol | mL | 9 | 0.42 | 21 | 2
20 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 49:2 |  | Ref | 1.8 | 1.12 | 0.68 | 2.92 | 0.56 | nmol | mL | 6 | 0.56 | 31 | 2
21 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 50:0 |  | Ref | 3.8 | 1.66 | 2.14 | 5.46 | 0.83 | nmol | mL | 11 | 0.83 | 22 | 2
22 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 50:1 |  | Ref | 38 | 20 | 18 | 58 | 10 | nmol | mL | 14 | 10 | 26 | 2
23 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 50:2 |  | Ref | 47 | 24 | 23 | 71 | 12 | nmol | mL | 15 | 12 | 26 | 2
24 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 50:3 |  | Ref | 23 | 13.2 | 9.8 | 36.2 | 6.6 | nmol | mL | 16 | 6.6 | 29 | 2
25 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 50:4 |  | Ref | 8.7 | 5.8 | 2.9 | 14.5 | 2.9 | nmol | mL | 15 | 2.9 | 34 | 2
26 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 50:5 |  | Ref | 1.6 | 1.28 | 0.32 | 2.88 | 0.64 | nmol | mL | 7 | 0.64 | 40 | 2
27 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 51:1 |  | Ref | 1.8 | 0.96 | 0.84 | 2.76 | 0.48 | nmol | mL | 7 | 0.48 | 27 | 2
28 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 51:2 |  | Ref | 4.8 | 2.2 | 2.6 | 7 | 1.1 | nmol | mL | 8 | 1.1 | 22 | 2
29 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 51:3 |  | Ref | 4.8 | 3.8 | 1 | 8.6 | 1.9 | nmol | mL | 5 | 1.9 | 39 | 2
30 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 52:1 |  | Ref | 14 | 5.8 | 8.2 | 19.8 | 2.9 | nmol | mL | 11 | 2.9 | 20 | 2
31 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 52:2 |  | Ref | 44 | 28 | 16 | 72 | 14 | nmol | mL | 16 | 14 | 33 | 2
32 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 52:3 |  | Ref | 100 | 58 | 42 | 158 | 29 | nmol | mL | 16 | 29 | 28 | 2
33 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 52:4 |  | Ref | 48 | 34 | 14 | 82 | 17 | nmol | mL | 15 | 17 | 35 | 2
34 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 52:5 |  | Ref | 15 | 11.4 | 3.6 | 26.4 | 5.7 | nmol | mL | 13 | 5.7 | 39 | 2
35 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 52:6 |  | Ref | 4 | 2.8 | 1.2 | 6.8 | 1.4 | nmol | mL | 8 | 1.4 | 35 | 2
36 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 52:7 |  | Ref | 0.39 | 0.26 | 0.13 | 0.65 | 0.13 | nmol | mL | 5 | 0.13 | 33 | 2
37 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 53:2 |  | Ref | 1.9 | 0.82 | 1.08 | 2.72 | 0.41 | nmol | mL | 9 | 0.41 | 21 | 2
38 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 53:3 |  | Ref | 3.7 | 2.2 | 1.5 | 5.9 | 1.1 | nmol | mL | 6 | 1.1 | 29 | 2
39 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 53:4 |  | Ref | 2.4 | 1.52 | 0.88 | 3.92 | 0.76 | nmol | mL | 6 | 0.76 | 32 | 2
40 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 54:1 |  | Ref | 3.2 | 1.82 | 1.38 | 5.02 | 0.91 | nmol | mL | 10 | 0.91 | 29 | 2
41 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 54:2 |  | Ref | 8.2 | 5.2 | 3 | 13.4 | 2.6 | nmol | mL | 13 | 2.6 | 31 | 2
42 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 54:3 |  | Ref | 26 | 19.6 | 6.4 | 45.6 | 9.8 | nmol | mL | 15 | 9.8 | 37 | 2
43 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 54:4 |  | Ref | 36 | 26 | 10 | 62 | 13 | nmol | mL | 15 | 13 | 35 | 2
44 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 54:5 |  | Ref | 27 | 22 | 5 | 49 | 11 | nmol | mL | 15 | 11 | 38 | 2
45 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 54:6 |  | Ref | 14 | 10.2 | 3.8 | 24.2 | 5.1 | nmol | mL | 16 | 5.1 | 37 | 2
46 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 54:7 |  | Ref | 5.6 | 3 | 2.6 | 8.6 | 1.5 | nmol | mL | 7 | 1.5 | 26 | 2
47 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 56:2 |  | Ref | 0.69 | 0.46 | 0.23 | 1.15 | 0.23 | nmol | mL | 5 | 0.23 | 33 | 2
48 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 56:3 |  | Ref | 1.4 | 0.28 | 1.12 | 1.68 | 0.14 | nmol | mL | 6 | 0.14 | 10 | 2
49 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 56:4 |  | Ref | 2 | 1.12 | 0.88 | 3.12 | 0.56 | nmol | mL | 10 | 0.56 | 28 | 2
50 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 56:5 |  | Ref | 4.1 | 2.8 | 1.3 | 6.9 | 1.4 | nmol | mL | 12 | 1.4 | 33 | 2
51 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 56:7 |  | Ref | 13 | 5.4 | 7.6 | 18.4 | 2.7 | nmol | mL | 8 | 2.7 | 20 | 2
52 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 56:9 |  | Ref | 0.71 | 0.54 | 0.17 | 1.25 | 0.27 | nmol | mL | 5 | 0.27 | 38 | 2
53 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 58:7 |  | Ref | 2 | 1.28 | 0.72 | 3.28 | 0.64 | nmol | mL | 5 | 0.64 | 32 | 2
54 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 58:8 |  | Ref | 0.68 | 0.42 | 0.26 | 1.1 | 0.21 | nmol | mL | 9 | 0.21 | 31 | 2
55 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 58:9 |  | Ref | 1.2 | 0.54 | 0.66 | 1.74 | 0.27 | nmol | mL | 6 | 0.27 | 22 | 2
56 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 14:0 |  | Ref | 1 | 0.4 | 0.6 | 1.4 | 0.2 | nmol | mL | 16 | 0.2 | 19 | 2
57 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 15:0 |  | Ref | 0.52 | 0.22 | 0.3 | 0.74 | 0.11 | nmol | mL | 9 | 0.11 | 22 | 2
58 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 16:0 |  | Ref | 73 | 22 | 51 | 95 | 11 | nmol | mL | 20 | 11 | 15 | 2
59 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC O-16:0 |  | Ref | 0.55 | 0.32 | 0.23 | 0.87 | 0.16 | nmol | mL | 10 | 0.16 | 29 | 2
60 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC P-16:0 |  | Ref | 0.46 | 0.26 | 0.2 | 0.72 | 0.13 | nmol | mL | 8 | 0.13 | 27 | 2
61 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 16:1 |  | Ref | 2.4 | 0.7 | 1.7 | 3.1 | 0.35 | nmol | mL | 19 | 0.35 | 15 | 2
62 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 17:0 |  | Ref | 1.4 | 0.48 | 0.92 | 1.88 | 0.24 | nmol | mL | 6 | 0.24 | 18 | 2
63 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 17:1 |  | Ref | 0.25 | 0.142 | 0.108 | 0.392 | 0.071 | nmol | mL | 6 | 0.071 | 29 | 2
64 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 18:0 |  | Ref | 27 | 6.6 | 20.4 | 33.6 | 3.3 | nmol | mL | 20 | 3.3 | 12 | 2
65 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC O-18:0 |  | Ref | 0.16 | 0.116 | 0.044 | 0.276 | 0.058 | nmol | mL | 6 | 0.058 | 36 | 2
66 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 18:1 |  | Ref | 18 | 4.6 | 13.4 | 22.6 | 2.3 | nmol | mL | 19 | 2.3 | 13 | 2
67 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 18:2 |  | Ref | 22 | 5.8 | 16.2 | 27.8 | 2.9 | nmol | mL | 19 | 2.9 | 13 | 2
68 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 18:3 |  | Ref | 0.44 | 0.26 | 0.18 | 0.7 | 0.13 | nmol | mL | 18 | 0.13 | 30 | 2
69 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 20:0 |  | Ref | 0.1 | 0.068 | 0.032 | 0.168 | 0.034 | nmol | mL | 7 | 0.034 | 34 | 2
70 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 20:1 |  | Ref | 0.19 | 0.048 | 0.142 | 0.238 | 0.024 | nmol | mL | 13 | 0.024 | 12 | 2
71 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 20:2 |  | Ref | 0.23 | 0.088 | 0.142 | 0.318 | 0.044 | nmol | mL | 9 | 0.044 | 19 | 2
72 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 20:3 |  | Ref | 1.8 | 0.52 | 1.28 | 2.32 | 0.26 | nmol | mL | 18 | 0.26 | 15 | 2
73 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 20:4 |  | Ref | 6 | 1.2 | 4.8 | 7.2 | 0.6 | nmol | mL | 20 | 0.6 | 10 | 2
74 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 20:5 |  | Ref | 0.33 | 0.184 | 0.146 | 0.514 | 0.092 | nmol | mL | 15 | 0.092 | 28 | 2
75 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 22:0 |  | Ref | 0.025 | 0.0034 | 0.0216 | 0.0284 | 0.0017 | nmol | mL | 5 | 0.0017 | 7 | 2
76 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 22:1 |  | Ref | 0.013 | 0.0092 | 0.0038 | 0.0222 | 0.0046 | nmol | mL | 5 | 0.0046 | 36 | 2
77 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 22:4 |  | Ref | 0.12 | 0.082 | 0.038 | 0.202 | 0.041 | nmol | mL | 8 | 0.041 | 33 | 2
78 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 22:5 |  | Ref | 0.43 | 0.26 | 0.17 | 0.69 | 0.13 | nmol | mL | 12 | 0.13 | 30 | 2
79 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 22:6 |  | Ref | 0.77 | 0.28 | 0.49 | 1.05 | 0.14 | nmol | mL | 17 | 0.14 | 18 | 2
80 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 24:0 |  | Ref | 0.046 | 0.03 | 0.016 | 0.076 | 0.015 | nmol | mL | 5 | 0.015 | 33 | 2
81 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPE 16:0 |  | Ref | 0.91 | 0.54 | 0.37 | 1.45 | 0.27 | nmol | mL | 14 | 0.27 | 29 | 2
82 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPE 18:0 |  | Ref | 1.6 | 1.1 | 0.5 | 2.7 | 0.55 | nmol | mL | 15 | 0.55 | 34 | 2
83 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPE 18:1 |  | Ref | 1.4 | 0.94 | 0.46 | 2.34 | 0.47 | nmol | mL | 14 | 0.47 | 35 | 2
84 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPE 18:2 |  | Ref | 1.9 | 1.12 | 0.78 | 3.02 | 0.56 | nmol | mL | 16 | 0.56 | 30 | 2
85 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPE 20:4 |  | Ref | 1.1 | 0.82 | 0.28 | 1.92 | 0.41 | nmol | mL | 14 | 0.41 | 37 | 2
86 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPE 22:6 |  | Ref | 0.52 | 0.36 | 0.16 | 0.88 | 0.18 | nmol | mL | 12 | 0.18 | 34 | 2
87 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 30:0 |  | Ref | 1.6 | 0.64 | 0.96 | 2.24 | 0.32 | nmol | mL | 11 | 0.32 | 20 | 2
88 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-30:0/29:0 |  | Ref | 0.072 | 0.052 | 0.02 | 0.124 | 0.026 | nmol | mL | 7 | 0.026 | 36 | 2
89 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-30:1/P-30:0 |  | Ref | 0.047 | 0.0192 | 0.0278 | 0.0662 | 0.0096 | nmol | mL | 7 | 0.0096 | 20 | 2
90 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 32:0 |  | Ref | 7.2 | 2 | 5.2 | 9.2 | 1 | nmol | mL | 18 | 1 | 14 | 2
91 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-32:0/31:0 |  | Ref | 1.5 | 0.82 | 0.68 | 2.32 | 0.41 | nmol | mL | 11 | 0.41 | 28 | 2
92 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 32:1 |  | Ref | 13 | 3.8 | 9.2 | 16.8 | 1.9 | nmol | mL | 18 | 1.9 | 15 | 2
93 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-32:1/P-32:0/31:1 |  | Ref | 1.6 | 0.48 | 1.12 | 2.08 | 0.24 | nmol | mL | 11 | 0.24 | 14 | 2
94 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-32:2/P-32:1/31:2 |  | Ref | 0.34 | 0.186 | 0.154 | 0.526 | 0.093 | nmol | mL | 8 | 0.093 | 28 | 2
95 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 32:3 |  | Ref | 0.42 | 0.28 | 0.14 | 0.7 | 0.14 | nmol | mL | 8 | 0.14 | 34 | 2
96 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC P-33:1/32:2 |  | Ref | 2.6 | 0.74 | 1.86 | 3.34 | 0.37 | nmol | mL | 16 | 0.37 | 14 | 2
97 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 34:0 |  | Ref | 2.1 | 0.74 | 1.36 | 2.84 | 0.37 | nmol | mL | 12 | 0.37 | 18 | 2
98 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-34:0/33:0 |  | Ref | 0.76 | 0.34 | 0.42 | 1.1 | 0.17 | nmol | mL | 10 | 0.17 | 22 | 2
99 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 34:1 |  | Ref | 120 | 42 | 78 | 162 | 21 | nmol | mL | 19 | 21 | 17 | 2
100 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-34:1/P-34:0/33:1 |  | Ref | 4.9 | 1.72 | 3.18 | 6.62 | 0.86 | nmol | mL | 17 | 0.86 | 17 | 2
101 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-34:2/P-34:1/33:2 |  | Ref | 5.2 | 2.6 | 2.6 | 7.8 | 1.3 | nmol | mL | 17 | 1.3 | 25 | 2
102 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-34:3/P-34:2/33:3 |  | Ref | 4.7 | 1.76 | 2.94 | 6.46 | 0.88 | nmol | mL | 12 | 0.88 | 19 | 2
103 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC P-35:1/34:2 |  | Ref | 240 | 94 | 146 | 334 | 47 | nmol | mL | 18 | 47 | 19 | 2
104 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC P-35:2/34:3 |  | Ref | 12 | 3.4 | 8.6 | 15.4 | 1.7 | nmol | mL | 18 | 1.7 | 14 | 2
105 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-35:4/34:4 |  | Ref | 1 | 0.5 | 0.5 | 1.5 | 0.25 | nmol | mL | 9 | 0.25 | 24 | 2
106 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 34:5 |  | Ref | 0.034 | 0.009 | 0.025 | 0.043 | 0.0045 | nmol | mL | 5 | 0.0045 | 13 | 2
107 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 36:1 |  | Ref | 26 | 9.2 | 16.8 | 35.2 | 4.6 | nmol | mL | 17 | 4.6 | 17 | 2
108 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-36:1/P-36:0/35:1 |  | Ref | 3.5 | 1.98 | 1.52 | 5.48 | 0.99 | nmol | mL | 16 | 0.99 | 28 | 2
109 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 36:2 |  | Ref | 140 | 50 | 90 | 190 | 25 | nmol | mL | 18 | 25 | 17 | 2
110 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-36:2/P-36:1/35:2 |  | Ref | 7.4 | 3.4 | 4 | 10.8 | 1.7 | nmol | mL | 17 | 1.7 | 22 | 2
111 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 36:3 |  | Ref | 100 | 28 | 72 | 128 | 14 | nmol | mL | 17 | 14 | 14 | 2
112 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-36:3/P-36:2/35:3 |  | Ref | 3.7 | 1.64 | 2.06 | 5.34 | 0.82 | nmol | mL | 12 | 0.82 | 22 | 2
113 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 36:4 |  | Ref | 150 | 56 | 94 | 206 | 28 | nmol | mL | 19 | 28 | 19 | 2
114 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-36:4/P-36:3/35:4 |  | Ref | 12 | 2.8 | 9.2 | 14.8 | 1.4 | nmol | mL | 17 | 1.4 | 12 | 2
115 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 36:5 |  | Ref | 11 | 3.6 | 7.4 | 14.6 | 1.8 | nmol | mL | 16 | 1.8 | 17 | 2
116 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-36:5/P-36:4/35:5 |  | Ref | 6.9 | 3.2 | 3.7 | 10.1 | 1.6 | nmol | mL | 11 | 1.6 | 23 | 2
117 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC P-36:5/35:6 |  | Ref | 0.3 | 0.188 | 0.112 | 0.488 | 0.094 | nmol | mL | 5 | 0.094 | 31 | 2
118 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 36:6 |  | Ref | 0.28 | 0.176 | 0.104 | 0.456 | 0.088 | nmol | mL | 8 | 0.088 | 32 | 2
119 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 38:2 |  | Ref | 2.3 | 0.4 | 1.9 | 2.7 | 0.2 | nmol | mL | 15 | 0.2 | 9 | 2
120 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-38:2/37:2 |  | Ref | 0.98 | 0.64 | 0.34 | 1.62 | 0.32 | nmol | mL | 6 | 0.32 | 32 | 2
121 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 38:3 |  | Ref | 26 | 10.4 | 15.6 | 36.4 | 5.2 | nmol | mL | 14 | 5.2 | 20 | 2
122 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-38:3/P-38:2/37:3 |  | Ref | 1.5 | 1.02 | 0.48 | 2.52 | 0.51 | nmol | mL | 14 | 0.51 | 34 | 2
123 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 38:4 |  | Ref | 84 | 28 | 56 | 112 | 14 | nmol | mL | 18 | 14 | 17 | 2
124 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-38:4/P-38:3/37:4 |  | Ref | 7.4 | 4 | 3.4 | 11.4 | 2 | nmol | mL | 12 | 2 | 27 | 2
125 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 38:5 |  | Ref | 42 | 15.8 | 26.2 | 57.8 | 7.9 | nmol | mL | 18 | 7.9 | 19 | 2
126 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-38:5/P-38:4/37:5 |  | Ref | 11 | 3.2 | 7.8 | 14.2 | 1.6 | nmol | mL | 16 | 1.6 | 14 | 2
127 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 38:6 |  | Ref | 41 | 8.8 | 32.2 | 49.8 | 4.4 | nmol | mL | 18 | 4.4 | 11 | 2
128 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-38:6/P-38:5/37:6 |  | Ref | 3.6 | 2 | 1.6 | 5.6 | 1 | nmol | mL | 12 | 1 | 29 | 2
129 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC P-38:6/36:0 |  | Ref | 1.2 | 0.78 | 0.42 | 1.98 | 0.39 | nmol | mL | 10 | 0.39 | 33 | 2
130 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 40:4 |  | Ref | 2.9 | 0.74 | 2.16 | 3.64 | 0.37 | nmol | mL | 18 | 0.37 | 13 | 2
131 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-40:2/P-40:1 |  | Ref | 0.069 | 0.042 | 0.027 | 0.111 | 0.021 | nmol | mL | 5 | 0.021 | 30 | 2
132 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-40:4/P-40:3/39:4 |  | Ref | 0.95 | 0.76 | 0.19 | 1.71 | 0.38 | nmol | mL | 8 | 0.38 | 40 | 2
133 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 40:5 |  | Ref | 6.7 | 2.2 | 4.5 | 8.9 | 1.1 | nmol | mL | 18 | 1.1 | 16 | 2
134 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-40:5/P-40:4/39:5 |  | Ref | 1.7 | 0.9 | 0.8 | 2.6 | 0.45 | nmol | mL | 12 | 0.45 | 27 | 2
135 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 40:6 |  | Ref | 14 | 5.2 | 8.8 | 19.2 | 2.6 | nmol | mL | 17 | 2.6 | 19 | 2
136 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 40:7 |  | Ref | 3.5 | 1.52 | 1.98 | 5.02 | 0.76 | nmol | mL | 16 | 0.26 | 21 | 2
137 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-40:7/P-40:6/39:7 |  | Ref | 1.1 | 0.46 | 0.64 | 1.56 | 0.23 | nmol | mL | 9 | 0.23 | 20 | 2
138 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 40:8 |  | Ref | 0.73 | 0.4 | 0.33 | 1.13 | 0.2 | nmol | mL | 14 | 0.2 | 28 | 2
139 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-42:5/P-42:4 |  | Ref | 0.79 | 0.24 | 0.55 | 1.03 | 0.12 | nmol | mL | 7 | 0.12 | 15 | 2
140 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 32:1 |  | Ref | 0.34 | 0.24 | 0.1 | 0.58 | 0.12 | nmol | mL | 6 | 0.12 | 36 | 2
141 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 34:1 |  | Ref | 1.2 | 0.34 | 0.86 | 1.54 | 0.17 | nmol | mL | 14 | 0.17 | 14 | 2
142 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 34:2 |  | Ref | 2.2 | 0.52 | 1.68 | 2.72 | 0.26 | nmol | mL | 16 | 0.26 | 12 | 2
143 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE O-34:2/P-34:1 |  | Ref | 0.78 | 0.34 | 0.44 | 1.12 | 0.17 | nmol | mL | 11 | 0.17 | 22 | 2
144 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE O-34:3/P-34:2 |  | Ref | 1.5 | 0.82 | 0.68 | 2.32 | 0.41 | nmol | mL | 11 | 0.41 | 27 | 2
145 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 36:0 |  | Ref | 0.28 | 0.2 | 0.08 | 0.48 | 0.1 | nmol | mL | 11 | 0.1 | 36 | 2
146 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 36:1 |  | Ref | 1.3 | 0.52 | 0.78 | 1.82 | 0.26 | nmol | mL | 14 | 0.26 | 20 | 2
147 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 36:2 |  | Ref | 6.7 | 1.58 | 5.12 | 8.28 | 0.79 | nmol | mL | 16 | 0.79 | 12 | 2
148 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE O-36:2/P-36:1/35:2 |  | Ref | 0.93 | 0.44 | 0.49 | 1.37 | 0.22 | nmol | mL | 12 | 0.22 | 23 | 2
149 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 36:3 |  | Ref | 2.4 | 0.76 | 1.64 | 3.16 | 0.38 | nmol | mL | 16 | 0.38 | 16 | 2
150 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE O-36:3/P-36:2/35:3 |  | Ref | 3.2 | 1.52 | 1.68 | 4.72 | 0.76 | nmol | mL | 15 | 0.76 | 24 | 2
151 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 36:4 |  | Ref | 3.1 | 0.78 | 2.32 | 3.88 | 0.39 | nmol | mL | 16 | 0.39 | 13 | 2
152 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE O-36:4/P-36:3 |  | Ref | 1.6 | 0.58 | 1.02 | 2.18 | 0.29 | nmol | mL | 14 | 0.29 | 18 | 2
153 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE O-36:5/P-36:4 |  | Ref | 4.9 | 3.8 | 1.1 | 8.7 | 1.9 | nmol | mL | 15 | 1.9 | 38 | 2
154 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 38:3 |  | Ref | 0.95 | 0.4 | 0.55 | 1.35 | 0.2 | nmol | mL | 14 | 0.2 | 21 | 2
155 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 38:4 |  | Ref | 8.1 | 2.4 | 5.7 | 10.5 | 1.2 | nmol | mL | 16 | 1.2 | 15 | 2
156 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE O-38:4/P-38:3/37:4 |  | Ref | 0.94 | 0.36 | 0.58 | 1.3 | 0.18 | nmol | mL | 9 | 0.18 | 19 | 2
157 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 38:5 |  | Ref | 2.7 | 0.94 | 1.76 | 3.64 | 0.47 | nmol | mL | 12 | 0.47 | 17 | 2
158 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE O-38:5/P-38:4 |  | Ref | 5.8 | 3.8 | 2 | 9.6 | 1.9 | nmol | mL | 17 | 1.9 | 33 | 2
159 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 38:6 |  | Ref | 3.2 | 1.18 | 2.02 | 4.38 | 0.59 | nmol | mL | 15 | 0.59 | 19 | 2
160 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE O-38:6/P-38:5 |  | Ref | 4.9 | 2.4 | 2.5 | 7.3 | 1.2 | nmol | mL | 16 | 1.2 | 25 | 2
161 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE O-38:7/P-38:6 |  | Ref | 3.5 | 1.96 | 1.54 | 5.46 | 0.98 | nmol | mL | 8 | 0.98 | 28 | 2
162 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 40:4 |  | Ref | 0.26 | 0.164 | 0.096 | 0.424 | 0.082 | nmol | mL | 10 | 0.082 | 31 | 2
163 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 40:5 |  | Ref | 0.73 | 0.46 | 0.27 | 1.19 | 0.23 | nmol | mL | 12 | 0.23 | 31 | 2
164 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE O-40:5/P-40:4/39:5 |  | Ref | 0.73 | 0.26 | 0.47 | 0.99 | 0.13 | nmol | mL | 12 | 0.13 | 17 | 2
165 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 40:6 |  | Ref | 1.8 | 0.72 | 1.08 | 2.52 | 0.36 | nmol | mL | 14 | 0.36 | 20 | 2
166 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE O-40:6/P-40:5/39:6 |  | Ref | 1.3 | 0.62 | 0.68 | 1.92 | 0.31 | nmol | mL | 14 | 0.31 | 23 | 2
167 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 40:7 |  | Ref | 0.77 | 0.52 | 0.25 | 1.29 | 0.26 | nmol | mL | 11 | 0.26 | 33 | 2
168 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE O-40:7/P-40:6/39:7 |  | Ref | 2.5 | 1.44 | 1.06 | 3.94 | 0.72 | nmol | mL | 14 | 0.72 | 29 | 2
169 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PI 32:1 |  | Ref | 0.56 | 0.22 | 0.34 | 0.78 | 0.11 | nmol | mL | 10 | 0.11 | 19 | 2
170 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PI 34:1 |  | Ref | 2.4 | 0.84 | 1.56 | 3.24 | 0.42 | nmol | mL | 14 | 0.42 | 17 | 2
171 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PI 34:2 |  | Ref | 2.8 | 0.76 | 2.04 | 3.56 | 0.38 | nmol | mL | 14 | 0.38 | 14 | 2
172 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PI 36:1 |  | Ref | 2.1 | 1.18 | 0.92 | 3.28 | 0.59 | nmol | mL | 13 | 0.59 | 28 | 2
173 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PI 36:2 |  | Ref | 7.7 | 1.86 | 5.84 | 9.56 | 0.93 | nmol | mL | 15 | 0.93 | 12 | 2
174 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PI 36:3 |  | Ref | 2.2 | 0.58 | 1.62 | 2.78 | 0.29 | nmol | mL | 14 | 0.29 | 14 | 2
175 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PI 36:4 |  | Ref | 3 | 0.96 | 2.04 | 3.96 | 0.48 | nmol | mL | 14 | 0.48 | 16 | 2
176 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PI 38:3 |  | Ref | 3.4 | 1.08 | 2.32 | 4.48 | 0.54 | nmol | mL | 14 | 0.54 | 16 | 2
177 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PI 38:4 |  | Ref | 19 | 4.4 | 14.6 | 23.4 | 2.2 | nmol | mL | 17 | 2.2 | 11 | 2
178 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PI 38:5 |  | Ref | 2.5 | 0.88 | 1.62 | 3.38 | 0.44 | nmol | mL | 15 | 0.44 | 18 | 2
179 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PI 38:6 |  | Ref | 0.32 | 0.062 | 0.258 | 0.382 | 0.031 | nmol | mL | 10 | 0.031 | 10 | 2
180 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PI 40:4 |  | Ref | 0.3 | 0.084 | 0.216 | 0.384 | 0.042 | nmol | mL | 7 | 0.042 | 14 | 2
181 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PI 40:6 |  | Ref | 0.84 | 0.32 | 0.52 | 1.16 | 0.16 | nmol | mL | 12 | 0.16 | 19 | 2
182 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PG 36:2 |  | Ref | 0.67 | 0.48 | 0.19 | 1.15 | 0.24 | nmol | mL | 6 | 0.24 | 36 | 2
183 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | HexCer d34:1 |  | Ref | 0.86 | 0.42 | 0.44 | 1.28 | 0.21 | nmol | mL | 6 | 0.21 | 25 | 2
184 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | HexCer d36:1 |  | Ref | 0.13 | 0.086 | 0.044 | 0.216 | 0.043 | nmol | mL | 5 | 0.043 | 34 | 2
185 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | HexCer d40:1 |  | Ref | 2.4 | 1.36 | 1.04 | 3.76 | 0.68 | nmol | mL | 5 | 0.68 | 28 | 2
186 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | HexCer d42:1 |  | Ref | 2.7 | 1.46 | 1.24 | 4.16 | 0.73 | nmol | mL | 6 | 0.73 | 27 | 2
187 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CER d34:1 |  | Ref | 0.28 | 0.088 | 0.192 | 0.368 | 0.044 | nmol | mL | 17 | 0.044 | 16 | 2
188 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CER d36:1 |  | Ref | 0.12 | 0.042 | 0.078 | 0.162 | 0.021 | nmol | mL | 14 | 0.021 | 17 | 2
189 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CER d38:1 |  | Ref | 0.11 | 0.042 | 0.068 | 0.152 | 0.021 | nmol | mL | 16 | 0.021 | 20 | 2
190 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CER d40:1 |  | Ref | 0.65 | 0.24 | 0.41 | 0.89 | 0.12 | nmol | mL | 18 | 0.12 | 18 | 2
191 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CER d40:2 |  | Ref | 0.15 | 0.042 | 0.108 | 0.192 | 0.021 | nmol | mL | 6 | 0.021 | 14 | 2
192 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CER d41:1 |  | Ref | 0.67 | 0.54 | 0.13 | 1.21 | 0.27 | nmol | mL | 7 | 0.27 | 40 | 2
193 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CER d42:1 |  | Ref | 1.9 | 0.94 | 0.96 | 2.84 | 0.47 | nmol | mL | 19 | 0.47 | 24 | 2
194 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CER d42:2 |  | Ref | 0.82 | 0.2 | 0.62 | 1.02 | 0.1 | nmol | mL | 19 | 0.1 | 12 | 2
195 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d31:1 |  | Ref | 0.19 | 0.098 | 0.092 | 0.288 | 0.049 | nmol | mL | 5 | 0.049 | 25 | 2
196 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d32:1 |  | Ref | 8.4 | 2.8 | 5.6 | 11.2 | 1.4 | nmol | mL | 14 | 1.4 | 17 | 2
197 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d32:2 |  | Ref | 0.66 | 0.48 | 0.18 | 1.14 | 0.24 | nmol | mL | 10 | 0.24 | 36 | 2
198 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d33:1 |  | Ref | 4.7 | 1.28 | 3.42 | 5.98 | 0.64 | nmol | mL | 14 | 0.64 | 14 | 2
199 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d34:0 |  | Ref | 5.8 | 2.6 | 3.2 | 8.4 | 1.3 | nmol | mL | 14 | 1.3 | 22 | 2
200 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d34:1 |  | Ref | 100 | 30 | 70 | 130 | 15 | nmol | mL | 21 | 15 | 15 | 2
201 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d34:2 |  | Ref | 16 | 4.4 | 11.6 | 20.4 | 2.2 | nmol | mL | 17 | 2.2 | 14 | 2
202 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d35:1 |  | Ref | 2.5 | 1.16 | 1.34 | 3.66 | 0.58 | nmol | mL | 9 | 0.58 | 23 | 2
203 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d35:2 |  | Ref | 0.52 | 0.42 | 0.1 | 0.94 | 0.21 | nmol | mL | 6 | 0.21 | 39 | 2
204 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d36:0 |  | Ref | 2 | 0.98 | 1.02 | 2.98 | 0.49 | nmol | mL | 11 | 0.49 | 24 | 2
205 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d36:1 |  | Ref | 20 | 7.4 | 12.6 | 27.4 | 3.7 | nmol | mL | 22 | 3.7 | 18 | 2
206 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d36:2 |  | Ref | 9.6 | 3 | 6.6 | 12.6 | 1.5 | nmol | mL | 22 | 1.5 | 16 | 2
207 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d36:3 |  | Ref | 1.3 | 0.82 | 0.48 | 2.12 | 0.41 | nmol | mL | 13 | 0.41 | 31 | 2
208 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d37:1 |  | Ref | 1 | 0.46 | 0.54 | 1.46 | 0.23 | nmol | mL | 11 | 0.23 | 23 | 2
209 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d38:1 |  | Ref | 11 | 6.2 | 4.8 | 17.2 | 3.1 | nmol | mL | 17 | 3.1 | 27 | 2
210 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d38:2 |  | Ref | 5.2 | 2.6 | 2.6 | 7.8 | 1.3 | nmol | mL | 17 | 1.3 | 25 | 2
211 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d38:3 |  | Ref | 0.61 | 0.48 | 0.13 | 1.09 | 0.24 | nmol | mL | 8 | 0.24 | 39 | 2
212 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d39:1 |  | Ref | 3.6 | 2 | 1.6 | 5.6 | 1 | nmol | mL | 14 | 1 | 29 | 2
213 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d39:2 |  | Ref | 0.61 | 0.32 | 0.29 | 0.93 | 0.16 | nmol | mL | 9 | 0.16 | 26 | 2
214 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d40:1 |  | Ref | 20 | 10.2 | 9.8 | 30.2 | 5.1 | nmol | mL | 17 | 5.1 | 25 | 2
215 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d40:2 |  | Ref | 12 | 5.6 | 6.4 | 17.6 | 2.8 | nmol | mL | 15 | 2.8 | 24 | 2
216 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d40:3 |  | Ref | 2.2 | 1.58 | 0.62 | 3.78 | 0.79 | nmol | mL | 8 | 0.79 | 37 | 2
217 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d41:1 |  | Ref | 7.7 | 4.2 | 3.5 | 11.9 | 2.1 | nmol | mL | 14 | 2.1 | 27 | 2
218 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d41:2 |  | Ref | 5.8 | 2.8 | 3 | 8.6 | 1.4 | nmol | mL | 14 | 1.4 | 24 | 2
219 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d41:3 |  | Ref | 0.77 | 0.6 | 0.17 | 1.37 | 0.3 | nmol | mL | 7 | 0.3 | 39 | 2
220 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d42:1 |  | Ref | 20 | 10.8 | 9.2 | 30.8 | 5.4 | nmol | mL | 21 | 5.4 | 28 | 2
221 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d42:2 |  | Ref | 44 | 22 | 22 | 66 | 11 | nmol | mL | 18 | 11 | 25 | 2
222 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d42:3 |  | Ref | 17 | 9.4 | 7.6 | 26.4 | 4.7 | nmol | mL | 12 | 4.7 | 27 | 2
223 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d43:2 |  | Ref | 1 | 0.58 | 0.42 | 1.58 | 0.29 | nmol | mL | 10 | 0.29 | 29 | 2
224 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d44:2 |  | Ref | 0.4 | 0.26 | 0.14 | 0.66 | 0.13 | nmol | mL | 9 | 0.13 | 32 | 2
225 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 14:0 |  | Ref | 16 | 12 | 4 | 28 | 6 | nmol | mL | 7 | 6 | 37 | 2
226 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 15:0 |  | Ref | 5.3 | 3.6 | 1.7 | 8.9 | 1.8 | nmol | mL | 6 | 1.8 | 34 | 2
227 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 16:0 |  | Ref | 210 | 116 | 94 | 326 | 58 | nmol | mL | 13 | 58 | 28 | 2
228 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 16:1 |  | Ref | 100 | 54 | 46 | 154 | 27 | nmol | mL | 11 | 27 | 27 | 2
229 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 16:2 |  | Ref | 1.9 | 0.92 | 0.98 | 2.82 | 0.46 | nmol | mL | 5 | 0.46 | 25 | 2
230 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 17:1 |  | Ref | 8.2 | 2 | 6.2 | 10.2 | 1 | nmol | mL | 9 | 1 | 13 | 2
231 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 18:0 |  | Ref | 15 | 7.4 | 7.6 | 22.4 | 3.7 | nmol | mL | 7 | 3.7 | 25 | 2
232 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 18:1 |  | Ref | 450 | 220 | 230 | 670 | 110 | nmol | mL | 14 | 110 | 25 | 2
233 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 18:2 |  | Ref | 1700 | 860 | 840 | 2560 | 430 | nmol | mL | 14 | 430 | 26 | 2
234 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 18:3 |  | Ref | 84 | 48 | 36 | 132 | 24 | nmol | mL | 13 | 24 | 28 | 2
235 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 20:3 |  | Ref | 35 | 24 | 11 | 59 | 12 | nmol | mL | 13 | 12 | 35 | 2
236 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 20:4 |  | Ref | 350 | 116 | 234 | 466 | 58 | nmol | mL | 14 | 58 | 17 | 2
237 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 20:5 |  | Ref | 38 | 17.2 | 20.8 | 55.2 | 8.6 | nmol | mL | 12 | 8.6 | 23 | 2
238 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 22:5 |  | Ref | 4.1 | 3.2 | 0.9 | 7.3 | 1.6 | nmol | mL | 6 | 1.6 | 39 | 2
239 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 22:6 |  | Ref | 37 | 19 | 18 | 56 | 9.5 | nmol | mL | 11 | 9.5 | 26 | 2
240 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | Cholesterol |  | Ref | 770 | 220 | 550 | 990 | 110 | nmol | mL | 8 | 110 | 14 | 2
241 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CDCA |  | Ref | 0.3 | 0.22 | 0.08 | 0.52 | 0.11 | nmol | mL | 7 | 0.11 | 38 | 2
242 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CA |  | Ref | 0.12 | 0.068 | 0.052 | 0.188 | 0.034 | nmol | mL | 9 | 0.034 | 28 | 2
243 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | DCA |  | Ref | 0.35 | 0.166 | 0.184 | 0.516 | 0.083 | nmol | mL | 9 | 0.083 | 24 | 2
244 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | GCDCA |  | Ref | 1.1 | 0.36 | 0.74 | 1.46 | 0.18 | nmol | mL | 8 | 0.18 | 17 | 2
245 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | GDCA |  | Ref | 0.43 | 0.138 | 0.292 | 0.568 | 0.069 | nmol | mL | 7 | 0.069 | 16 | 2
246 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | GLCA |  | Ref | 0.025 | 0.0036 | 0.0214 | 0.0286 | 0.0018 | nmol | mL | 6 | 0.0018 | 7 | 2
247 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | GUDCA |  | Ref | 0.15 | 0.048 | 0.102 | 0.198 | 0.024 | nmol | mL | 6 | 0.024 | 16 | 2
248 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | GCA |  | Ref | 0.24 | 0.138 | 0.102 | 0.378 | 0.069 | nmol | mL | 6 | 0.069 | 29 | 2
249 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LCA |  | Ref | 0.014 | 0.0072 | 0.0068 | 0.0212 | 0.0036 | nmol | mL | 8 | 0.0036 | 26 | 2
250 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TCDCA |  | Ref | 0.084 | 0.01 | 0.074 | 0.094 | 0.005 | nmol | mL | 9 | 0.005 | 6 | 2
251 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TCA |  | Ref | 0.026 | 0.0112 | 0.0148 | 0.0372 | 0.0056 | nmol | mL | 9 | 0.0056 | 22 | 2
252 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TDCA |  | Ref | 0.04 | 0.0128 | 0.0272 | 0.0528 | 0.0064 | nmol | mL | 8 | 0.0064 | 16 | 2
253 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TLCA |  | Ref | 0.0027 | 0.00138 | 0.00132 | 0.00408 | 0.00069 | nmol | mL | 5 | 0.00069 | 26 | 2
254 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | UDCA |  | Ref | 0.11 | 0.048 | 0.062 | 0.158 | 0.024 | nmol | mL | 8 | 0.024 | 22 | 2
255 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TLCA-S |  | Inf | 0.088 | 0.036 | 0.052 | 0.124 | 0.018 | nmol | mL | 3 | 0.018 | 21 | 2
256 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | MCA |  | Inf | 0.0056 | 0.00026 | 0.00534 | 0.00586 | 0.00013 | nmol | mL | 3 | 0.00013 | 2 | 2
257 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 22:0 |  | Inf | 0.35 | 0.068 | 0.282 | 0.418 | 0.034 | nmol | mL | 4 | 0.034 | 10 | 2
258 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 24:0 |  | Inf | 0.28 | 0.104 | 0.176 | 0.384 | 0.052 | nmol | mL | 3 | 0.052 | 18 | 2
259 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CE 24:1 |  | Inf | 0.15 | 0.028 | 0.122 | 0.178 | 0.014 | nmol | mL | 3 | 0.014 | 9 | 2
260 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | Total Cholesterol |  | Inf | 4000 | 48 | 3952 | 4048 | 24 | nmol | mL | 3 | 24 | 1 | 2
261 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | HexCer 38:1 |  | Inf | 0.2 | 0.0003 | 0.1997 | 0.2003 | 0.00015 | nmol | mL | 3 | 0.00015 | 0 | 2
262 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CER d16:0 |  | Inf | 0.052 | 0.015 | 0.037 | 0.067 | 0.0075 | nmol | mL | 3 | 0.0075 | 14 | 2
263 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CER d20:0 |  | Inf | 0.028 | 0.0152 | 0.0128 | 0.0432 | 0.0076 | nmol | mL | 3 | 0.0076 | 28 | 2
264 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CER d24:0 |  | Inf | 0.09 | 0.066 | 0.024 | 0.156 | 0.033 | nmol | mL | 3 | 0.033 | 37 | 2
265 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CER d39:1 |  | Inf | 0.093 | 0.032 | 0.061 | 0.125 | 0.016 | nmol | mL | 4 | 0.016 | 18 | 2
266 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CerOH d16:0 |  | Inf | 5.3 | 1.34 | 3.96 | 6.64 | 0.67 | nmol | mL | 3 | 0.67 | 13 | 2
267 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CerOH d22:0 |  | Inf | 0.26 | 0.092 | 0.168 | 0.352 | 0.046 | nmol | mL | 3 | 0.046 | 18 | 2
268 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CerOH d24:0 |  | Inf | 0.4 | 0.1 | 0.3 | 0.5 | 0.05 | nmol | mL | 3 | 0.05 | 12 | 2
269 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | CerOH d24:1 |  | Inf | 0.97 | 0.28 | 0.69 | 1.25 | 0.14 | nmol | mL | 3 | 0.14 | 15 | 2
270 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | DAG 40:4 |  | Inf | 0.91 | 0.68 | 0.23 | 1.59 | 0.34 | nmol | mL | 4 | 0.34 | 38 | 2
271 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | FFA 17:0 |  | Inf | 1.7 | 0.46 | 1.24 | 2.16 | 0.23 | nmol | mL | 3 | 0.23 | 14 | 2
272 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | FFA 17:1 |  | Inf | 0.86 | 0.3 | 0.56 | 1.16 | 0.15 | nmol | mL | 3 | 0.15 | 18 | 2
273 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | FFA 20:1 |  | Inf | 1.8 | 0.34 | 1.46 | 2.14 | 0.17 | nmol | mL | 3 | 0.17 | 9 | 2
274 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC O-18:1 |  | Inf | 0.41 | 0.26 | 0.15 | 0.67 | 0.13 | nmol | mL | 3 | 0.13 | 31 | 2
275 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC O-20:0 |  | Inf | 0.023 | 0.0098 | 0.0132 | 0.0328 | 0.0049 | nmol | mL | 4 | 0.0049 | 22 | 2
276 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 24:1 |  | Inf | 0.022 | 0.0142 | 0.0078 | 0.0362 | 0.0071 | nmol | mL | 3 | 0.0071 | 33 | 2
277 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | LPC 26:0 |  | Inf | 0.015 | 0.0062 | 0.0088 | 0.0212 | 0.0031 | nmol | mL | 3 | 0.0031 | 21 | 2
278 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC 28:0 |  | Inf | 0.16 | 0.05 | 0.11 | 0.21 | 0.025 | nmol | mL | 4 | 0.025 | 15 | 2
279 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-42:6/P-42:5 |  | Inf | 0.46 | 0.28 | 0.18 | 0.74 | 0.14 | nmol | mL | 4 | 0.14 | 29 | 2
280 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-42:7 |  | Inf | 0.067 | 0.0022 | 0.0648 | 0.0692 | 0.0011 | nmol | mL | 3 | 0.0011 | 2 | 2
281 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PC O-44:5/P-44:4 |  | Inf | 1.3 | 0.6 | 0.7 | 1.9 | 0.3 | nmol | mL | 3 | 0.3 | 24 | 2
282 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 34:3 |  | Inf | 0.14 | 0.04 | 0.1 | 0.18 | 0.02 | nmol | mL | 4 | 0.02 | 14 | 2
283 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PE 35:1 |  | Inf | 0.15 | 0.09 | 0.06 | 0.24 | 0.045 | nmol | mL | 4 | 0.045 | 30 | 2
284 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PG 34:2 |  | Inf | 0.55 | 0.44 | 0.11 | 0.99 | 0.22 | nmol | mL | 4 | 0.22 | 40 | 2
285 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PG 36:3 |  | Inf | 0.5 | 0.22 | 0.28 | 0.72 | 0.11 | nmol | mL | 4 | 0.11 | 22 | 2
286 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PS 38:1 |  | Inf | 0.23 | 0.092 | 0.138 | 0.322 | 0.046 | nmol | mL | 3 | 0.046 | 20 | 2
287 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PS 38:2 |  | Inf | 0.24 | 0.08 | 0.16 | 0.32 | 0.04 | nmol | mL | 3 | 0.04 | 16 | 2
288 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d35:0 |  | Inf | 0.044 | 0.0144 | 0.0296 | 0.0584 | 0.0072 | nmol | mL | 3 | 0.0072 | 16 | 2
289 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | SM d42:5 |  | Inf | 0.36 | 0.042 | 0.318 | 0.402 | 0.021 | nmol | mL | 3 | 0.021 | 6 | 2
290 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | dhS1P |  | Inf | 0.1 | 0.07 | 0.03 | 0.17 | 0.035 | nmol | mL | 3 | 0.035 | 35 | 2
291 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | S1P |  | Inf | 0.42 | 0.0152 | 0.4048 | 0.4352 | 0.0076 | nmol | mL | 3 | 0.0076 | 2 | 2
292 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 44:3 |  | Inf | 0.18 | 0.0188 | 0.1612 | 0.1988 | 0.0094 | nmol | mL | 4 | 0.0094 | 5 | 2
293 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 47:2 |  | Inf | 0.21 | 0.054 | 0.156 | 0.264 | 0.027 | nmol | mL | 3 | 0.027 | 13 | 2
294 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 49:0 |  | Inf | 0.31 | 0.11 | 0.2 | 0.42 | 0.055 | nmol | mL | 3 | 0.055 | 18 | 2
295 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 54:8 |  | Inf | 0.92 | 0.6 | 0.32 | 1.52 | 0.3 | nmol | mL | 4 | 0.3 | 32 | 2
296 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 55:3 |  | Inf | 0.43 | 0.3 | 0.13 | 0.73 | 0.15 | nmol | mL | 3 | 0.15 | 34 | 2
297 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 58:11 |  | Inf | 0.11 | 0.036 | 0.074 | 0.146 | 0.018 | nmol | mL | 3 | 0.018 | 16 | 2
298 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | TAG 58:3 |  | Inf | 0.19 | 0.038 | 0.152 | 0.228 | 0.019 | nmol | mL | 3 | 0.019 | 10 | 2
299 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 11,12-DiHETrE |  | Inf | 0.00082 | 0.00056 | 0.00026 | 0.00138 | 0.00028 | nmol | mL | 3 | 0.00028 | 34 | 2
300 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 11-HDoHE |  | Inf | 0.00063 | 0.000184 | 0.000446 | 0.000814 | 9.2e-05 | nmol | mL | 3 | 9.2e-05 | 15 | 2
301 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 12,13-DiHOME |  | Inf | 0.0051 | 0.00076 | 0.00434 | 0.00586 | 0.00038 | nmol | mL | 3 | 0.00038 | 7 | 2
302 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 12,13-EpOME |  | Inf | 0.0069 | 0.004 | 0.0029 | 0.0109 | 0.002 | nmol | mL | 3 | 0.002 | 28 | 2
303 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 12-HHTrE |  | Inf | 0.00023 | 0.000106 | 0.000124 | 0.000336 | 5.3e-05 | nmol | mL | 3 | 5.3e-05 | 23 | 2
304 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 13-HOTrE |  | Inf | 0.00054 | 0.000112 | 0.000428 | 0.000652 | 5.6e-05 | nmol | mL | 3 | 5.6e-05 | 10 | 2
305 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 14-HDoHE |  | Inf | 0.0013 | 0.00022 | 0.00108 | 0.00152 | 0.00011 | nmol | mL | 4 | 0.00011 | 8 | 2
306 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 17-HDoHE |  | Inf | 0.00082 | 7.2e-05 | 0.000748 | 0.000892 | 3.6e-05 | nmol | mL | 3 | 3.6e-05 | 4 | 2
307 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 18-HEPE |  | Inf | 0.00028 | 0.000138 | 0.000142 | 0.000418 | 6.9e-05 | nmol | mL | 3 | 6.9e-05 | 25 | 2
308 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 20-HETE |  | Inf | 0.0021 | 0.00106 | 0.00104 | 0.00316 | 0.00053 | nmol | mL | 3 | 0.00053 | 25 | 2
309 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 5,6-EET |  | Inf | 0.00082 | 0.00056 | 0.00026 | 0.00138 | 0.00028 | nmol | mL | 3 | 0.00028 | 34 | 2
310 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 5-HEPE |  | Inf | 0.00085 | 3.2e-05 | 0.000818 | 0.000882 | 1.6e-05 | nmol | mL | 4 | 1.6e-05 | 2 | 2
311 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 8-HETE |  | Inf | 0.00098 | 0.00044 | 0.00054 | 0.00142 | 0.00022 | nmol | mL | 4 | 0.00022 | 22 | 2
312 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 8-HETrE |  | Inf | 0.00038 | 0.000186 | 0.000194 | 0.000566 | 9.3e-05 | nmol | mL | 3 | 9.3e-05 | 24 | 2
313 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 9,10-DiHOME |  | Inf | 0.0067 | 0.00088 | 0.00582 | 0.00758 | 0.00044 | nmol | mL | 3 | 0.00044 | 7 | 2
314 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 9-HEPE |  | Inf | 0.00043 | 0.000174 | 0.000256 | 0.000604 | 8.7e-05 | nmol | mL | 4 | 8.7e-05 | 20 | 2
315 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 9-HETE |  | Inf | 0.00085 | 0.000164 | 0.000686 | 0.001014 | 8.2e-05 | nmol | mL | 4 | 8.2e-05 | 10 | 2
316 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 9-HODE |  | Inf | 0.01 | 0.0052 | 0.0048 | 0.0152 | 0.0026 | nmol | mL | 3 | 0.0026 | 25 | 2
317 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | 9-OxoODE |  | Inf | 0.0073 | 0.0026 | 0.0047 | 0.0099 | 0.0013 | nmol | mL | 3 | 0.0013 | 18 | 2
318 | SRM 1950 - "Metabolites in Frozen Human Plasma" | Metabolites in Frozen Human Plasma | 2016-04-21 | PGE2 |  | Inf | 3.5e-05 | 2.8e-05 | 7e-06 | 6.3e-05 | 1.4e-05 | nmol | mL | 4 | 1.4e-05 | 40 | 2

## Data Overview (hidden)
Concentrations (nmol/mL)
Class | Lipid Name | Sum Composition

## Summed Data (hidden)
Concentrations

## Formulas (hidden)
One or two tails | TG | Sphingolipids |  | ±
0 | 0 | #VALUE! | 0 | ** Consensus mean ± standard uncertainty.
* Measurement mean ± 1 standard deviation.